## Sheet1
Make | Meril | Meril | Horiba | TBM | Aspen | Sysmex | Sysmex | AGD | Medonic | Nihon Kodan
Model | CelQuant Edge | CelQuant 3i | ABX Micros ES 60 | H360 | AD 3200 | XP-300 | XQ-320 | HT 340 | M32 | Celltac α MEK-1301 | Meril Benefits | Remarks
Throughput Open Tube (OT): | 60 samples/hour | 60 samples/hour | 60 samples/hour | 60 samples/hour | 60 samples/hour | 60 samples/hour | 70 samples/hour | Results within 45 seconds | 60 samples/hour | 60 samples/hour | Multiple language option
Throughput Close Tube (CT): | NA | NA | 50 samples/hour | NA | NA | NA | NA | NA | ≥ 45 Samples | NA (Available with MEK-1302) | Noiseless
Operating temperature: | 10 ~ 30° C | 10 ~ 35° C | (16-30° C)* | 15-30°C | 15-30°C | 15 °C TO 30 °C | 15 °C TO 30 °C | 15 C - 35 C | 18 - 32°C | 15 - 30ºC | tracking of reagent consumptions by real time RFID systems to track inventory
WB Specimen Volume: | ≤14 µL | 9 µL (venous or capillary mode) | 10 µL | Whole Blood : 9 μL | Venous: 9 μL Venous Capillary: 9 μL capillary | Whole Blood Mode: 50μL | Whole Blood Mode: 16µL | Whole Blood Mode: 7.5 μL Dilution/High mode: 3.75 μL | Open tube: ≤ 110 µL (uses 22.5 µL for analysis) Cap Piercer: ≤ 250 µL (uses 22.5 µL for analysis) Autoloader: ≤ 300 µL (uses 22.5 µL for analysis) | Normal mode: CBC + WBC 3 part differential 20 μL Predilution mode: CBC 10 or 20 μL Capillary mode: CBC 20 μL (Optional | Automatically re counts the sample after unclogging process without recollecting the blood sample. Can avoid wastage of blood sample and save time. | Can we incorporate capillary mode?
Pre-Dilute Specimen Volume | 20 µL (pre-dilute mode) | 20 µL (pre-dilute mode) | Not Specified | 20 µL (Pre-dilute or Capillary) | Pre-diluted: 20 μL capillary | Not Specified | 65 μL (pre-diluted mode, requires 20 μL blood) | 3.75 µL (Dilution or High mode) | Not Specified | 10 or 20 µL (predilution mode) | Back flushing, soaking and burning, automated unclog process before & after each treatment.
Calibration mode | Fresh blood, Manual, Automated | Fresh blood, Manual, Automated | Fresh blood, Manual, Automated | Manual Calibration Calibration using a Calibrator (such as ELite H Cal) Calibration using Fresh Blood | Manual calibration & auto calibration | SCS-1000 calibrator material; accessible via menu | SCS-1000 calibrator; supports manual and automatic calibration | Manual & Automatic calibration | Manual & Automatic calibration | Automated, supports QC lot management up to 25, Westgard multirule; barcode reader for QC registration | edge can be more ideal for clinics and small hospital as well.
Counting modes(Whole blood, pre dilute, capillary) | Whole blood and pre dilute. Capillary not available | Whole blood and pre dilute. Capillary not available | Whole blood and pre dilute. Capillary available | Whole Blood, Pre-dilute & Capillary available | Venous whole blood, whole blood, pre-diluted, capillary available | Whole Blood (WB), Pre-diluted (PD) | Whole blood, pre-dilute | Normal mode (whole blood), Dilution or High mode | Whole blood, pre-dilute, capillary available | Whole blood, pre-dilute, capillary available | Easy maintenance and minimal hands on time.
WBC Detection Principles:-
Method | Electrical Impedance | Electrical Impedance | Electrical Impedance | Electrical Impedance | Electrical Impedance | Direct Current detection method | Direct current method | Electrical Impedance | Electrical impedance | Electrical Impedance (DynaHelix Flow technology)
Ruby aperture diameter | 100 µm | 100 µm | 80 µm | Not Specified | 100 μm | Not Specified | Not Specified | Not Specified | 80 µm | Not Specified
RBC/PLT Detection Principles:-
Method | Electrical Impedance | Electrical Impedance | Electrical Impedance | Electrical Impedance | Electrical Impedance | Direct Current detection method | Direct current method | Electrical Impedance | Electrical impedance | Electrical impedance (DynaHelix Flow technology)
Ruby aperture diameter | 70 µm | 70 µm | 50 µm | Not Specified | 68 µm | Not Specified | Not Specified | Not Specified | 50 µm | Not Specified
HGB Measurement:-
Method | Photometry, Cyanide Free Colorimetry | Photometry, Cyanide Free Colorimetry | Photometry | Cyanide Free Colorimetry | Cyanide Free Colorimetry | Non-cyanide method | Non-cyanide HGB analysis method | Non-cyanide Photometric Hemoglobin detection | Cyanide Free Colorimetry | Colorimetric method
Wavelength | 540nm | standard: 530–540 nm | 550 nm | 630 nm | typically 530 nm | Not Specified | spectrophotometric at 555 nm | Not Specified | 570 nm | typicallly: 555 nm
HCT/MCV Measurement:-
Method | HCT: Calculated MCV: Measured | HCT: Calculated MCV: Measured | Numeric integration | HCT: Calculated MCV: Measured | HCT: Calculated MCV: Directly Measured | Cumulative pulse height detection method | Calculated from RBC count and volume (impedance method) | Calculated from RBC histogram | Calculated from RBC histogram | HCT: Calculated from RBC histogram MCV: Calculated |  | Can we add calculation for MCV?
PARAMETERS & PERFORMANCE DATA
Reportable Parameters | 21 Parameters : WBC, RBC, HGB, HCT, MCV MCH, MCHC, RDW-CV, RDW-SD, PLT, MPV, PCT, PDW, P-LCR, P-LCC, Lymph%/#, Mid %/#, Gran%/# | 21 Parameters : WBC, Lym#, Mid#, Gran#, Lym%, Mid%, Gran%, RBC, HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD, PLT, MPV, PDW, PCT, HGB, P-LCC, P-LCR | 19 Parameters : WBC, RBC, RDW CV, HGB, MPV, HCT, MCV, PLT, MCH, MCHC, LYM# & LYM%, MON# & MON%, GRA# & GRA%, PDW*, PCT*, RDW SD*. | 24 parameters : WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD, PLT, MPV, PCT, PDW-CV, PDW-SD, P-LCR, P-LCC, LYM%, MID%, GRAN%, LYM#, MID#, GRAN#, Mentzer index, RDWI | 22 parameters : WBC, RBC, HGB, PLT, LYM%, MID%, GRAN%, LYM#, MID#, GRAN#, HCT, MCV, MC+, MCHC, RDW-SD, RDW-CV, MPV, PDW, PCT | WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT, LYM#, LYM%, MXD#, MXD%, NEUT#, NEUT%, RDW-SD, RDW-CV, MPV | 22 parameters : WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT, LYM%, MXD%, NEUT%, LYM#, MXD#, NEUT#, RDW-SD, RDW-CV, PDW, MPV, P-LCR, PCT | 22 Parameters : WBC, LY, MID, GR, LY%, MID%, GR%, Hgb, RBC, HCT, MCV, MCH, MCHC, RDW–SD, RDW–CV, PLT, PCT, MPV, PDW-SD, PDW-CV, Mentzer index, Sehgal index 4 Research parameters: PLCR*, PLCC*, GLR*, PLR* | 22 Parameters : WBC, LYM, MID, GRAN, LYM%, MID%, GRAN%, RBC, MCV, HCT, PLT, MPV, HGB, MCH, MCHC, RDW%, RDW, PCT, PDW%, PDW, P-LCR, P-LCC | 20 Parametrs : WBC, LY%, MO%, GR%, LY#, MO#, GR#, RBC, HGB, HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD, PLT, PCT, MPV, PDW, P-LCR 4 Research parameters: P-LCC, Mentzer Index, RDWI, NLR
Research Parameters | NA | PLR, NLR | NA | Mentzer index, RDWI, P-LCR, P-LCC | NA | NA | NA | PLCR*, PLCC*, GLR*, PLR* | NA | P-LCC, Mentzer Index, RDWI, NLR |  | Can we incorporate additional research parameters?
Linearity:
WBC (103 /mm3) | 0 - 100 | 0 - 100 | 100 - 150 | 0 - 300 | Not Specified | 1.0 - 99.9 X 103/μL | Not Specified | 0.0 ~ 100.0 | 0.20 - 130.0 x 10 /L | 0.20 to 99.9 x 10 raised to 3 /μl
RBC (106 /mm3) | 0 - 7 | 0 - 7 | 8 - 18 x | 0.00 - 8.50 | Not Specified | 0.30 - 7.00 X 106/μL | Not Specified | 0.00 ~ 10 | 0.30 - 7.00 x 10 /L | 0.02 to 8.00 x 10 raised to 6 /μl
Hgb (g/dl) | 0 - 25 | 0 - 25 | 26 - 30 | 0 - 25.0 | Not Specified | 0.1 - 25.0 g/dL | Not Specified | 0.00 ~ 30 | 2.0 - 24.0 g/dL | 0.10 to 25.0 g/dL
HCT/MCV | MCV: 0 - 250 | MCV: 0 - 250 | HCT: 80 - 90% | HCT: 0 - 67 | Not Specified | 10 - 60% | Not Specified | Not Specified | MCV: NA | 20.0 to 60.0%
PLT | 0 - 3000 | 0 - 3000 | 2200 - 6000 x 103 /mm3 | 0 – 3000 | Not Specified | 10 - 999 X 103/μL | Not Specified | 0 ~ 1600 | 10 - 1800 x 10 /L PLT Conc. Mode: 10 - 5000 x 10 /L | 10.0 to 1490.0 x 10 3 /μL
Precision:
WBC | ≤ 2.0% | ≤ 2.0% | <2.5 | Not Specified | 0.00 ~ 300.00) x 10 raised 9 /L | Not Specified | Not Specified | < 3.00% | ≤ 1.6 % | 2.0% or less
RBC | ≤ 1.5% | ≤ 1.5% | <2.0 | Not Specified | 0.00 ~ 8.50) x 10 12 /L | Not Specified | Not Specified | < 2.00% | ≤ 0.8 % | 1.5% or less
Hgb | ≤ 1.5% | ≤ 1.5% | <1.5 | Not Specified | 0 ~ 250g/L | Not Specified | Not Specified | < 2.00% | ≤ 0.6 % | 1.5% or less
HCT/MCV | ≤ 0.5% | ≤ 0.5% | HCT: <2.0 | Not Specified | HCT: 0 ~ 67% | Not Specified | Not Specified | HCT: < 2.00% | MCV: ≤ 0.5 % | HCT: 1.5% or less MCV: 1.0% or less
PLT | ≤ 5.0% | ≤ 5.0% | <5.0 | Not Specified | 0 ~ 3000x10 9 /L | Not Specified | Not Specified | < 5.00% | ≤ 2.9 % | 4.0% or less
QC Data | includes 9 QC files, L-J Graph, Automatically draw quality | Supports daily QC, L-J charting, multi-level controls | Levey-Jennings graphics XB management | L-J, X-BAR | Multilevel GC setup with L-J, X, SD, CV | Mean, SD.CV, Levey- Jennings plots and X-B plots | Levey-Jennings; X-barM, tri-level QC product for all parameters | Tri-level QC with LJ plot | Mean, SD.CV, Levey- Jennings plots and X-B plots | Not Specified
Display | 10.4" LCD Touchscreen | LCD, 640×480 resolution | 8.4" LCD Touchscreen | 10.4" LCD Touchscreen | 10.4" LCD Touchscreen | Color LCD touchscreen | 10.4” color LCD touchscreen | 8“ Touch Panel Colour LCD with back light | 7 inch WVGA true color(24-bit) touch screen | Color LCD touchscreen
Automatic clog removal facilities | Flush and Enhanced Flush functions for aperture clog removal | Flush and Enhanced Flush functions for aperture clog removal | Yes | Prevent build-up with anti-clog technology | Yes | Automatic cleaning and rinse cycles; monthly transducer cleaning | Not specified | Not explicitly specified | Yes | Not explicitly specified; DynaHelix Flow reduces clogging risk |  | DynaHelix flow technology
Automatic cleaning of probe | Rinse, Enhanced Cleaning, and maintenance programs | Rinse, Enhanced Cleaning, and maintenance programs | Yes, after every test | Not Specified | Yes | Not Specified | Not specified | Not explicitly specified | Yes | Not explicitly specified; regular cleaning with recommended reagents
Data storage | 50000 with histograms | Upto 50000 with histograms | Upto 1000 (expandable upto 50000) samples with histograms | Upto 50000 with 3 histograms (RBC, WBC, PLT) | Upto 50,000 sample result with histograms | Upto 40,000 results with histograms | Upto 500,000 results with histograms | Upto 20000 with 3 Histograms (RBC, WBC and Platelets) | Upto 50,000 samples | Upto 50,000 data with histogram
Ports(USB, LAN, RS232) | 4 USB, 1 LAN, 1 RS232 | 4 USB, 2 RS232, Printer port | 2 USB, 2 RS232, 1 PS2, 1 RJ45 | 4 USB, 1 LAN Port | 4 USB ports & bi-directional LIS connection | USB, LAN (Ethernet), RS232, barcode reader port | USB ×5, LAN ×2, RS232, external printer port, barcode reader port | RS232, USB | USB, LAN, RS232 | 1 LAN, 2 USB, 3 RS-232C
Thermal in built printer | Yes, 80 mm | Yes | Yes, 58 mm | Yes, 57 mm | Yes | Yes, 57 mm | No (external printer optional) | Yes | Yes | Not built-in; external printer supported
External printer | Yes | Yes | Yes | Yes | Yes | Yes | Supported (optional) | Yes | Yes | Supported (optional)
CERTIFICATION
CE - IVD, US - FDA, | CE-IVD, ISO | CE-IVD, ISO | CE-IVD, ISO | CE-IVD; US FDA not specified | CE-IVD, ISO | CE-IVD, US FDA 510(k) cleared | CE-IVD | CE - IVD | CE-IVD, FDA 510(k) cleared | CE-IVD certified; US FDA status not specified
Physical Specification
Dimensions: | 368 x 448 x 475 mm (l x w x h) | 420 x 295 x 430 mm (l x w x h) | 430 x 360 x 360 mm (h x w x d) | 364 x 477 x 417(mm) (l x w x h) | 364 x 417 x 477 mm (l x w x h) | 419.1 X 482.6 X 355.6 mm (w x d x h) | 355.6 x 444.5 x 431.8 mm (w x d x h) | 430 x 280 x 425 mm (l x w x h) | 395 x 295 x 475 mm (l x w x h) | 230 x 450 x 428 mm (w x d x h)
Weight OT: | 23 kg | 28 kg | 14 kg | 25 kg | ≤ 25 kg | Approx 30 kg | Approx 22 kg | Approx 17 kg | ≤ 18 kg | 20 kg
Weight CT: | NA | NA | 17 kg |  |  |  | NA | NA | NA | NA
Multiple language option
Noiseless
Note on G6: Sayali: 14ul of blood sample should be sufficient for testing
Note on H7: Sayali: 20ul of Pre diluted blood should be sufficient for testing.
Note on D12: Sayali: Should have atleast 100um to ensure accurate counting and sizing of cells as they pass through.
Note on J12: Sayali: Should have atleast 100um to ensure accurate counting and sizing of cells as they pass through.
Note on D15: Sayali: The aperture diameter should be 70 um for better discrimination of small cells.
Note on J15: Sayali: The aperture diameter should be 70 um for better discrimination of small cells.
Note on D37: Sayali: Screen size should be 10" or more
Note on I37: Sayali: Screen size should be 10" or more
Note on J37: Sayali: Screen size should be 10" or more
Note on G40: Sayali: Storage should be upto 50,000 samples with atleast 3 histograms
Note on I40: Sayali: Storage should be upto 50,000 samples with atleast 3 histograms
Note on D41: Sayali: Should have atleast 4 ports for Keyboard, mouse, disk & Printer. Should have bi directional LIS connection
Note on G41: Sayali: Should have atleast 4 ports for Keyboard, mouse, disk & Printer. Should have bi directional LIS connection
Note on I41: Sayali: Should have atleast 4 ports for Keyboard, mouse, disk & Printer. Should have bi directional LIS connection
Note on K41: Sayali: Should have atleast 4 ports for Keyboard, mouse, disk & Printer. Should have bi directional LIS connection
Note on D42: Sayali: Should have atleast 80mm built in printer which provides full report with histogram. Should have optional printer interface.
Note on E42: Sayali: Should have atleast 80mm built in printer which provides full report with histogram. Should have optional printer interface.
Note on G42: Sayali: Should have atleast 80mm built in printer which provides full report with histogram. Should have optional printer interface.
Note on K42: Sayali: Should have atleast 80mm built in printer which provides full report with histogram. Should have optional printer interface.
Note on D47: Sayali: More storage space.